### Phenotypic Detection of (-Lactamase Resistance in Gram-Negative Bacilli: Testing and Interpretation Guide (Rev. 3-17-10)

Paul C. Schreckenberger, Ph.D., Violeta Rekasius, MT(ASCP)

INTRODUCTION:

To detect isolates with AmpC, K1, and KPC enzymes, as well as those with ESBLs, it is desirable to test all Enterobacteriaceae with a combination of antibiotics that will allow detection of resistant mechanisms. In our laboratory we have expanded the ESBL confirmatory disk test to include 10 antibiotic disks and have applied the test to all members of the Enterobacteriaceae that have a susceptibility pattern that is suspicious for the presence of an ESBL, AmpC, K1 or KPC type resistance gene. Through the application of this test we have been able to successfully detect antibiotic resistance in many species of Enterobacteriaceae that would have not been detected using our automated susceptibility testing system.

WHEN TO PERFORM 10-DISK TEST:

Set up 10-disk test to confirm presence of ESBL in following instances:

- Any E. coli or Klebsiella when phenotype does not agree with ESBL confirmation test on Vitek, MicroScan, or other automated susceptibility system (e.g. ceftazidime is I or R but ESBL confirmatory test is Neg)

- Any Enterobacteriaceae when one of the 3rd gen. cephalosporins tests I or R and no ESBL confirmatory test result available

- Any Enterobacteriaceae when atypical or multi-drug resistant pattern exists (e.g. P. mirabilis resistant to multiple drugs)

- Any Enterobacteriaceae resistant to all drugs except imipenem

SPECIMEN:

The specimen consists of a pure isolate of the Enterobacteriaceae, which has a susceptibility result that is consistent with an ESBL or an AMPC pattern, ie.; ceftazidime is I/R, cefrtriaxone is I/R, aztreonam is I/R and requires confirmation by a disk method.

MATERIALS:

- Antibiotic disks are placed in 12 cartridge dispenser, kept in fridge (2-8C), until use:

Aztreonam (30)

Ceftazidime (30)

Ceftazidime + clavulante (30/10)

Cefotaxime (30)

Cefotaxime + clavulante (30/10)

Cefoxitin (30)

Ceftriaxone (30)

Cefepime (30)

Ertapenem (10)

Imipenem(10)

- Mueller Hinton (MH) agar plate, 150 mm , kept in fridge (2-8C), until use

- sterile saline or tryptic soy broth (TSB)

- sterile swabs

- 0.5 McFarland barium turbidity standard / photometer (colorimeter)

METHOD:

- Allow the MH agar plate and disk dispenser to come to room temperature before use.

- Prepare a 0.5 McFarland standard of the organism to be tested in sterile saline or TSB.  Standardize the inoculum using the colorimeter.

- Streak the bacterial suspension evenly in 3 planes onto the surface of the MH agar plate, using a cotton swab.  Rim the edge of the plate.

- Place the disk dispenser over the MH agar plate and depress the knob.  This will allow the antibiotic disks to dispense and automatically “tamp” the disk into place.

- All of the disks must be placed on the same MH agar plate in a specified order (See Figure 1)

- Incubate the MH agar plate overnight in a non-CO2 incubator at 35C.

- The following day read and record all zones of inhibition.

RESULTS:

- Detection of ESBLs (ceftazidime and cefotaxime disks with and without clavulanic acid are used to detect ESBLs)

A.  If the zone size increases 5 mm or more when clavulanate is added compared to the drug alone the isolate is considered an ESBL.  Only one antibiotic must be "reversed" by the clavulanate to be an ESBL.  Note: For MYSPACE organisms if the zone size increases 3 mm or more with clavulanic-containing disk compared to the drug alone then the isolate is considered an ESBL. See reference 14.

CAZ/CLA – 22 mm C	CAZ – 11 mm

22 –11 = > 5mm

= ESBL

B.  If an “enhancement” or extension of the zone of inhibition is seen between any of the cephalosporin antibiotics and the clavulanate containing disks, the presence of an ESBL can be predicted. This phenomenon is often referred to as the “KEYHOLE” effect, or “CLAVULANIC” effect and is indicative of ESBL production.

- Detection of AmpC beta lactamases (cefepime and cefoxitin disks are used to detect ampC beta lactamases)

  - AmpC strains are resistant to the cephamycins (ie; cefoxitin and cefotetan).

  - AmpC strains are susceptible to cefepime.

  - High level AmpC producers causes resistance to all 1st, 2nd and 3rd generation cephalosporins, the beta lactam-inhibitor drugs and the monobactams (ie; aztreonam).

- Detection of K1 beta lactamases (aztreonam, ceftazidime, cefotaxime and ceftriaxone disks are used to detect K1 beta lactamases)

Cefoxitin S = Neg AmpC

No Clavulanic Effect =

Neg ESBL

Aztreonam - R Ceftazidime - S

Cefotaxime S:   Ceftriaxone R

- Detection of Carbapenemase (ertapenem and imipenem disks are used to screen for carbapenemase resistance)

                - Imipenem - S

Ertapenem - R

                - Suggests possible KPC which should be confirmed with Hodge test or sent to reference lab for confirmation

- Record all disk diffusion mm zone size readings in the culture work up.

- If ESBL is confirmed, change/override any previous susceptibility result to resistant, if the antibiotic is a penicillin, cephalosporin, or monobactam regardless of how the drug tests, following CLSI interpretive guidelines for ESBL.  Refer to CLSI document M100-S18. Report cephamycins (ie; cefoxitin) and beta lactam inhibitor drugs as they test (in other words report as susceptible if they test susceptible, do not override).

- If ESBL is not confirmed then report drugs as they test. For example if organism is shown to be an ampC or K1, report drugs as they test, do not override and make resistant

- If ESBL is present along with ampC or K1 then apply the ESBL reporting rules and report all penicillins, cephalosporins and monobactams as resistant.

- If KPC is confirmed then report all beta lactam drugs as resistance regardless of how they test

QUALITY CONTROL:

Disk diffusion testing is performed weekly with ATCC# 700603 Klebsiella pneumoniae and E. coli ATCC 25922 following CLSI guidelines.  If correct quality control results are not obtained, the test is invalid and patient results cannot be reported.

REFERENCE:

Clinical and Laboratory Standards Institute. Performance Standards for Antimicrobial Susceptibility Testing, Nineteenth  Informational Supplement.  CLSI document M100-S19. Clinical and Laboratory Standards Institute, 940 West Valley Road, Suite 1400, Wayne, Pennsylvania 19087-1898 USA, 2009.

## General Review Articles

- Livermore DM, Winstanley TB, Shannon KP. Interpretative reading: recognizing the unusual and inferring resistance mechanisms from resistance phenotypes.
J Antimicrob Chemother. 2001 Jul;48 Suppl 1:87-102

- Thomson KS, Moland ES. Version 2000: the new (-lactamases of Gram-negative bacteria at the dawn of the new millennium (Review). Microbes and Infection, 2000;2:1225-1235

Review of Phenotypic Detection Methods

- Drieux L, Brossier F, Sougakoff W, Jarlier V. Phenotypic detection of extended-spectrum (-lactamase production in Enterobacteriaceae: review and bench guide. Clin Microbiol Infect 2008; 14 (Suppl. 1) 90-103.

- Livermore DM, Brown DF. Detection of beta-lactamase-mediated resistance. J Antimicrob Chemother. 2001 Jul;48 Suppl 1:59-64.

## ESBL Review Articles

- Bradford PA. Extended-spectrum beta-lactamases in the 21st century: characterization, epidemiology, and detection of this important resistance threat. Clin Microbiol Rev. 2001 Oct;14(4):933-51, table of contents. Review.

- Jacoby GA, Munoz-Price LS: The new beta-lactamases.
N Engl J Med. 2005 Jan 27;352(4):380-91. Review

- Paterson DL, Bonomo RA: Extended-spectrum beta-lactamases: a clinical update.
Clin Microbiol Rev. 2005 Oct;18(4):657-86. Review.

- Schreckenberger, P.C. Conundrums in the Laboratory Detection of Antimicrobial Resistant Gram-Negative Bacteria. Reviews in Medical Microbiology; 15:45-49, April 2004.

- Thomson KS: Controversies about extended-spectrum and AmpC beta-lactamases. Emerg Infect Dis. 2001 Mar-Apr;7(2):333-6. Review.

## AmpC B-Lactamases Review Articles

- Jacoby GA. AmpC (-Lactamases. Clin Microbiol Rev; 2009 Jan; 22(1):161-182.

## Double-Disk Potentiation Method

- Coudron PE, Moland ES, Sanders CC: Occurrence and detection of extended-spectrum beta-lactamases in members of the family Enterobacteriaceae at a veterans medical center: seek and you may find. J Clin Microbiol. 1997 Oct;35(10):2593-7.

- Jarlier V, Nicolas MH, Fournier G, Philippon A: Extended broad-spectrum beta-lactamases conferring transferable resistance to newer beta-lactam agents in Enterobacteriaceae: hospital prevalence and susceptibility patterns. Rev Infect Dis. 1988 Jul-Aug;10(4):867-78.

- Pitout JD, Reisbig MD, Venter EC, Church DL, Hanson ND. Modification of the double-disk test for detection of enterobacteriaceae producing extended-spectrum and AmpC beta-lactamases. J Clin Microbiol. 2003 Aug;41(8):3933-5.

- Towne TG, Lewis JS 2nd, Herrera M, Wickes B, Jorgensen JH: Detection of SHV-type extended-spectrum beta-lactamase in Enterobacter isolates. J Clin Microbiol. 2010 Jan;48(1):298-9.

- Tzelepi E, Giakkoupi P, Sofianou D, Loukova V, Kemeroglou A, Tsakris A: Detection of extended-spectrum beta-lactamases in clinical isolates of Enterobacter cloacae and Enterobacter aerogenes. J Clin Microbiol. 2000 Feb;38(2):542-6.

- Yan JJ, Hsueh PR, Lu JJ, Chang FY, Shyr JM, Wan JH, Liu YC, Chuang YC, Yang YC, Tsao SM, Wu HH, Wang LS, Lin TP, Wu HM, Chen HM, Wu JJ: Extended-spectrum beta-lactamases and plasmid-mediated AmpC enzymes among clinical isolates of Escherichia coli and Klebsiella pneumoniae from seven medical centers in Taiwan. Antimicrob Agents Chemother. 2006 May;50(5):1861-4.

- Yan JJ, Ko WC, Wu HM, Tsai SH, Chuang CL, Wu JJ. Complexity of Klebsiella pneumoniae isolates resistant to both cephamycins and extended-spectrum cephalosporins at a teaching hospital in Taiwan. J Clin Microbiol. 2004 Nov;42(11):5337-40.

## Hodge Test

- Anderson KF, Lonsway DR, Rasheed JK, Biddle J, Jensen B, McDougal LK, Carey RB, Thompson A, Stocker S, Limbago B, Patel JB. Evaluation of methods to identify the Klebsiella pneumoniae carbapenemase in Enterobacteriaceae. J Clin Microbiol. 2007 Aug;45(8):2723-5.

- Lee K, Chong Y, Shin HB, Kim YA, Yong D, Yum JH. Modified Hodge and EDTA-disk synergy tests to screen metallo-beta-lactamase-producing strains of Pseudomonas and Acinetobacter species. Clin Microbiol Infect. 2001 Feb;7(2):88-91.

- Lee K, Lim YS, Yong D, Yum JH, Chong Y. Evaluation of the Hodge test and the imipenem-EDTA double-disk synergy test for differentiating metallo-beta-lactamase-producing isolates of Pseudomonas spp. and Acinetobacter spp. J Clin Microbiol. 2003 Oct;41(10):4623-9

# Text boxes

Double Disk Potentiation Method – K. oxytoca with chromosomal K1 beta lactamase, ESBL Negative, AmpC Negative

22 mm

11mm

# Double Disk Potentiation Method – P. mirabilis with ESBL

Double Disk Potentiation Method - E. coli with plasmid mediated AmpC, ESBL Negative

##### No Clavulanic Effect

Cefepime (FEP) – S

Cefoxitin (FOX) – R

= AmpC

# Combination Disk (CLISI) Method – E.coli with  ESBL

##### Keyhole Formation between cefepime and clavulanic containing disk

##### = ESBL

Cefepime (FEP) – S

Cefoxitin (FOX) – R

= AmpC

Fig 1. Template for Disk Potentiation Method for Detecting

ESBL and ampC beta-lactamases

1

2

CTX-CLA

9

10

ATM

IMP

3

8

CTX

11

ERT

FEP

12

CAZ-CLA

4

CAZ

7

FOX

CRO

5

6

empty

empty

###### Abbreviation KEY

- cefotaxime-clavulanate

- aztreonam

- cefepime

- ceftriaxone

- empty

- empty

- cefoxitin

- ertapenem

- imipenem

10	cefotaxime

11	ceftazidime-clavulanate

12	ceftazidime

Double Disk Potentiation Method - E. cloacae with chromosomal AmpC, ESBL Positive

Double Disk Potentiation Method – K. pneumoniae with KPC beta lactamase

##### Keyhole Formation Around Clavulanic Containing Disk

##### = ESBL